PRIGo IIIB Calibration

Instruction Manual

Version:	1.2

Last Edit:	August 18, 2015

Author:	Andreas Isenegger, Bitcontrol GmbH

Filename:	PRIGO_CalibrationInstructions_1.2.odt

Document Revision History

Version	Date	Author	Changes

1.2	18-Aug-2015	A. Isenegger	Typos fixed, chapters 7, 8, 9 updated
1.1	29-Jun-2015	A. Isenegger	Chapters 8 and 9 extended, Author corrected
1.0	25-Jun-2015	A. Isenegger	Initial version

Table 1: Document revision history

# Summary
## Calibration from Scratch
A calibration from scratch includes all necessary steps to calibrate a Prigo after its assembly. It is required also if a linear stage has been replaced. Follow these steps in this order for a new stage:
1. Configure a Linear Stage in the MCS
2. Remove the Tetrahedron from the Prigo
3. Replace a Linear Stage
4. Configure the Zero Position of a Linear Stage
5. Attach the Tetrahedron to the Prigo
6. Optimize the IK Model Parameters
7. Optimize the Active Correction Parameters

## Calibration after Relocation
If an assembled Prigo is relocated from one Aerotech base stage to another, only the step Optimize the Active Correction Parameters has to be executed.

# Configure a Linear Stage in the MCS
Wear gloves when touching Smaract stages. See their user manual.
## Required Tools
- Smaract Hand Controller (green hand held box)
- Smaract MCS Configuration Tool (Windows application)

## Procedure
How to prepare a linear stage:
- If more than one stage has to be replaced, mark them temporarily with S1, S2, S3 or S4, as desired
- Stop the prigoTCP service
- Switch the MCS controller off
- Connect the stage to the control system on the channel it will finally be connected to
- Connect the Smaract hand controller
- Switch the MCS controller on
- Calibrate the linear stage
- Check if the stage moves and travels as expected
- Check the direction of the stage when moving positive; invert its direction if necessary using the Smaract MCS configuration tool:
  - Disconnect the Prigo PC from the MCS controller
  - Connect the Windows PC running the Smaract MCS Configuration Tool to the MCS controller
  - Start the Smaract MCS Configuration Tool (a Windows application) and change the direction of positive motion
  - Disconnect the Windows PC running the Smaract MCS Configuration Tool from the MCS controller
  - Connect the Prigo PC to the MCS controller
- Switch the MCS controller off
- Disconnect the stage from the control system
- Cover its bottom with a silicon film

# Remove the Tetrahedron from the Prigo
Wear gloves when touching Smaract stages. See their user manual.
## Assumptions
- The Prigo is mounted to an Aerotech rotation stage
- The rotation stage can rotate freely; this simplifies the manipulations

## Procedure
- Manipulate gently!
- Switch the MCS controller off
- At the front of the Prigo, disconnect the 2-wire cable to the magnet
- Loosen the 2 screws on the swing that hold the ball joint of the S4 stage's push rod and move the ball out of its bearing
- While holding the tetrahedron and the push rod with one hand, unlink the 4 springs that hold the tetrahedron on the 3 ball joints with the other hand; start with S2 and S3, unlink S1 last
- Gently lay the tetrahedron on a soft tissue below the Prigo
- If you need to completely remove the tetrahedron from the Prigo, unplug the motor cable of the phi rotary stage from the distribution box and remove all tape and cable ties that guide this cable to the phi stage. Then pull the connector through the center of the rotation stage.

# Replace a Linear Stage
Wear gloves when touching Smaract stages. See their user manual.
## Required Tools
- Standard Allen keys
- Long Allen key 1.5mm; ask Roman Schneider for the set of long Allen keys

## Procedure
How to replace a linear stage:
- Switch the MCS controller off
- Disconnect the Lemo connector from the distribution box
- If stage S1, S2 or S3 has to be mounted or dismounted, dismount stage S4 before
- Remove the flexor from the stage
- Loosen the 2 grub screws
- Loosen the screws that hold the stage on the black frame; use the long Allen key 1.5mm:
  - S1, S2, S3 are hold in place by 6 screws
  - S4 is hold in place by 4 screws; the 2 screws towards the front of the Prigo may not be inserted because the holes on the linear stage are missing
- Remove the old stage
- Mount the flexor to the new stage
- Mount the new stage:
- Optional, for simple Aerotech rotation stages only: Feed the cable through the center hole of the rotation stage
- For S1, S2, S3:
  - S4 needs to be removed
  - Mount the stage using 6 M2 x 8mm stainless steel screws with cylindrical Allen heads; don't tighten the screws yet, just almost
  - Center the stage longitudinally as good as possible: there is about 0.7mm play
  - Tighten the 2 grub screws slightly
  - Tighten the 6 M2 x 8mm stainless steel screws
  - Loosen the 2 grub screws
  - Tighten the 6 M2 x 8mm stainless steel screws again (not more force, just check)
  - Tighten the 2 grub screws
- For S4:
  - Mount the stage using 4 M2 x 5mm stainless steel screws with cylindrical Allen heads;
    Note: use the 4 screw holes close to the base of the Prigo; the 2 holes towards the tip of the Prigo don't fit! The mounting threads with 90mm distance on the bottom of the linear stage are not provided anymore by Smaract.
    Don't tighten the screws yet, just almost
  - Center the stage longitudinally as good as possible: there is about 0.7mm play
  - Tighten the 4 M2 x 5mm stainless steel screws
- Connect the stage to the distribution box

# Configure the Zero Position of a Linear Stage
Wear gloves when touching Smaract stages. See their user manual.
## Required Tools
- Prigo Zero Position Calibration tool
- Smaract Hand Controller (green hand held box)

## Procedure
- Stop the prigoTCP service
- Switch the MCS controller off
- Connect the Smaract hand controller
- Switch the MCS controller on
  Don't move any of the stages using the hand controller; this would activate position control for the moved axis, which is undesired at this point
- Search the reference mark on the linear stage
- For S1, S2, S3:
  - Remove the front most of the 4 screws that fixes the Flexor to the linear stage
  - For S2, S3: Remove the screw on the Flexor that holds the spring
  - Insert the two 1.5mm thick aluminum shims to fix the Flexor's tip
  - Mount the Prigo Zero Position Calibration tool
  - Move the linear stage by hand such that the Flexor's ball touches the calibration tool's front block
  - Set the offset of the linear stage to zero
  - Search the reference mark on the linear stage
  - Using the hand controller, move the stage 5mm away in positive direction
  - Power cycle the MCS controller
  - Search the reference mark on the linear stage
  - Using the hand controller, move the stage to its zero position; the ball should be touching the calibration tool's front block, with a tolerance of +-2um
    If the error is larger, repeat this procedure.
- For S4:
  - Shift the linear stage to its center position manually; center it visually as accurately as possible; maximum allowed error is +-100um
  - Set the offset of the linear stage to zero
  - Search the reference mark on the linear stage
  - Using the hand controller, move the stage 5mm away in positive direction
  - Power cycle the MCS controller
  - Search the reference mark on the linear stage
  - Using the hand controller, move the stage to its zero position; check its position visually; it shouldn't be off center more than +-100um
    If the error is larger, repeat this procedure.

# Attach the Tetrahedron to the Prigo
Assumptions:
- The Prigo is mounted to an Aerotech rotation stage
- The rotation stage can rotate freely; this simplifies the manipulations

Procedure:
- Manipulate gently!
- Switch the MCS controller off
- Shift S4 about 30mm back (towards the base of the Prigo)
- If the push rod has been removed from the swing: Loosen the 2 screws on the swing that hold the ball joint of the S4 stage's push rod and insert the push rod's ball into its bearing; tighten the 2 screws
- Hold the tetrahedron with one hand and push it gently against the 3 ball joints of the Flexors
- Link the 4 springs that hold the tetrahedron on the 3 ball joints with the other hand; start with S1, link S2 and S3 last
- At the front of the Prigo, connect the 2-wire cable to the magnet
- If the cable of the phi stage has been unplugged from the distribution box and removed, feed the cable through the center of the rotation stage and connect it to the distribution box; tape it to the black base body of the Prigo where necessary

# Optimize the IK Model Parameters
The parts and the assembly of the Prigo is afflicted with small mechanical deviations. Although they are individually smaller than 100um, they can sum up to a sphere of confusion of 500um radius when Omega is rotated.
In earlier procedures, many of the IK model parameters were included in the optimization process, and some of them were optimized until they got values far beyond their realistic tolerances.
In the current approach, only the parameters offsetS1, offsetS2, offsetS3, offsetS4, l11 and l44 have been changed. The goal is to achieve an SoC radius of 5um when Omega is rotated, and an SoC deviation of maximum 300um for a 90 degree Chi movement. Once this has been achieved, active correction is used to eliminate the remaining errors.
## Procedure
- Start the prigoTCP service; be patient, this takes a minute
- Using the medm PRIGO Control Panel:
  - Connect to the IOC
  - Initialize the Prigo's axes (move the distance sensors away beforehand, and back into position afterwards; the Z-axis is well suited for this purpose)
  - Check the Prigo applications and EPICS
  - Move the tip to the position
    - [SHX, SHY, SHZ] = [0mm, 0mm, 20mm]
    - [OX, OY, OZ] = [0mm, 0mm, 275mm]
- Switch active correction off (see 9.4.4)
- Set all IK parameters to their default values (see 9.4.2)
- Disengage the Omega axis
- Place the Capacitive Distance Sensors for low sensitivity measurements
- Setup the capacitive distance measurement system, connect its analog outputs with the related EPICS IOC, and for convenience with an oscilloscope that is placed nearby
- Set the sensitivity on the CPL290 capacitive distance measurement unit of all channels to LOW (62500nm/V)
- Mount the the 17.5mm long steel cylinder (the one that has a 2mm hole, 2 dots and a carved line on its side) carrying a 12.7mm (1/2 inch) steel sphere to the magnet; align the carved line on the magnet with the carved line on the steel cylinder; be precise
- Rotate the Omega axis by hand; using the linear tables move each sensor to the position where its signal amplitude is symmetrical around 0 when rotating Omega 360deg
- Correct the possibly large mounting tolerances of the magnet and the steel cylinder:
  - Position to Omega axis to 0deg manually
  - Run the IDL script phimeto.pro (see 9.3.7) with these parameters:
    IDL> phimeto, DATACOLLECTION=0, SENS=62500, EFAC=10000
    This centers the steel sphere relative to the Phi axis by changing the SH position and therefore reduces the excentricity of the sphere.
- Rotate the Omega axis by hand and change the offsetS4 parameter (see 9.4.1) until the magnet and cylinder align well with the Omega axis (rotate manually, check visually, compare with fix point in the background)
- This step shouldn't be necessary since the positions of the stages S1, S2 and S3 should be very precise (within 10um tolerance). Their offset on the MCS has been adjusted using the Zero Position Calibration tool (see chapter 5.).
  Change the offsets only within 50um maximum. If you still have a large deviation on the tip, look for other sources of the error, or repeat the zero position calibration (chapter 5.).

  Rotate the Omega axis by hand and change the offsetS1, offsetS2, offsetS3 parameters (see 9.4.1) until the SoC radius becomes smaller than 50um (check on the scope). If the value can't be reached, proceed to the next step.
- Repeat the previous two steps until the SoC radius comes in the range of 50um (check on the scope)
- Execute a Chi sweep, 0deg .. 90deg:
  - Switch the Active Correction off (see 9.4.4)
  - Run the IDL script phimeto.pro (see 9.3.7) with these parameters:
    IDL> phimeto, DATACOLLECTION=0, SENS=62500, EFAC=2000
  - Run the IDL script gometo.pro (see 9.3.6) with these parameters:
    IDL> gometo, DATACOLLECTION=0, SENS=62500, EFAC=500
  - Run the SoC application for recording and displaying the SoC of the Chi sweep

    OR
  - Run the IDL script chimeto.pro (see ) with these parameters:
    IDL> chimeto, SENS=62500, FILENAME='2015-08-03_01', MNUMBER=3
    (change the file name to your convenience)
  - Use the Matlab script evaluate.m to display the deviations graphically (see 8.1, paragraph Chi fit parameters how to do this)
- Check if the deviations are within this tolerance (procedure see below):
  - X-axis deviation: ±10um
  - Y-axis deviation: +300um
  - Z-axis deviation: ±300um
- If the deviations exceed the above values, change l44 by 100um. Observe the magnet's orientation relative to the horizon. It should stay horizontally. If it doesn't compensate by changing offsetS1 to offsetS3. Then repeat the Chi sweep.
- If it got better try to improve it by another change of 100um in the same direction. If it got worse, change l44 100um in the other direction and repeat the Chi sweep.
- If the steps above stop improving the deviations, try varying the parameter l11. It is the length of the magnet, its attachment to the rotation stage, and part of the rotation stage. This value might in reality have more than 100um deviation from the drawings and can therefore play a significant role in Chi's deviations.
- Once satisfied, store the new values in the prigoIK.xmp parameter file (see 9.4.1)

Finally check the SoC for Chi and Omega sweeps of the mechanically calibrated Prigo system. The values should be comparable to the values listed below, otherwise something went wrong in the previous steps:
- Omega rotation, -390deg .. +390deg, high sensitivity on distance sensors:
  - X-axis radius: ±5um
  - Y-axis radius: ±5um
  - Z-axis radius: ±2.5um
- Chi sweep, 0deg .. 90deg, low sensitivity on distance sensors:
  - X-axis deviation: ±10um
  - Y-axis deviation: +300um
  - Z-axis deviation: ±300um

Procedure how to perform this check:
- Switch the Active Correction off (see 9.4.4)
- Run the IDL script phimeto.pro (see 9.3.7) with these parameters:
  IDL> phimeto, DATACOLLECTION=0, SENS=62500, EFAC=2000
- Run the IDL script gometo.pro (see 9.3.6) with these parameters:
  IDL> gometo, DATACOLLECTION=0, SENS=62500, EFAC=500
- Run the SoC application for recording and displaying the SoC of Chi and Omega rotations
# Optimize the Active Correction Parameters
These parameters compensate for the sag induced by gravity when rotating Omega and the deviations between the mechanics (the reality) and the mathematical IK model. The software module that calculates these compensations is called Calibration.
## Procedure
- Start the prigoTCP service; be patient, this takes a minute
- Using the medm PRIGO Control Panel:
  - Connect to the IOC
  - Initialize the Prigo's axes (move the distance sensors away beforehand, and back into position afterwards; the Z-axis is well suited for this purpose)
  - Check the Prigo applications and EPICS
  - Move the tip to the position
    - [SHX, SHY, SHZ] = [0mm, 0mm, 20mm]
    - [OX, OY, OZ] = [0mm, 0mm, 275mm]
    - Omega = Phi = Chi = 0deg
- Engage the Omega axis
- Set the Prigo active correction fit parameters to 0
- Switch Active Correction on (see 9.4.4)
- Place the Capacitive Distance Sensors for low sensitivity measurements
- Run the IDL script phimeto.pro (see 9.3.7) with these parameters:
  IDL> phimeto, DATACOLLECTION=0, SENS=62500, EFAC=500

Omega fit parameters
- Run the IDL script gometo.pro (see 9.3.6) with these parameters:
  IDL> gometo, SENS=62500, EFAC=250, STARTANGLE=-390, ENDANGLE=390, MNUMBER=5, VELOCITY=2, FILENAME='2015-05-25_01' # Note the single quotes! Adapt the date.
- Run matlab and generate the coefficients of the fit functions for the Omega rotation:
  - $> cd data
  - $> matlab &
  - Edit the script evaluate.m and adapt the name of the data file and the number of data sets to the configuration applied to gometo.pro in the previous step:
    - for ii = 1:5
    - dset=load(sprintf('2015-05-25_01_%d.txt',ii));
  - Run the script evaluate.m; this loads the data from the data file into Matlab and displays the sphere of confusion graphically in two variants:
    >> evaluate
  - Run the module cftool; this is a curve fitting tool provided by MathWorks; it uses the data set loaded by evaluate.m
    >> cftool
  - Configure cftool:
    - X data: omega
    - Y data: X for a fit on the beamline X-axis; use Y data: Y for a fit on the Y-axis or Y data: Z for a fit on the Z-axis respectively (do one after the other)
    - Use Sum of Sine for the Omega corrections
    - Select 1 to 4 in Number of Terms:, depending on what achieves the best result without generating useless calculations inside Orchestra: when the RMSE value in the result section of the cftool window doesn't change much when a higher number of terms is selected, then the optimum of calculation efficiency is reached
- Copy the active correction fit parameters from cftool to Orchestra
- Go back to the step Run the IDL script gometo.pro and repeat the procedure until here with minor differences:
  - Run the IDL script gometo.pro (see 9.3.6) with these parameters:
    IDL> gometo, SENS=62500, EFAC=250, STARTANGLE=-390, ENDANGLE=390, MNUMBER=5, VELOCITY=2, FILENAME='2015-05-25_02'
  - If Matlab is still running, execute the command clear before running evaluate
  - Edit the script evaluate.m and adapt the name of the data file and the number of data sets to the configuration applied to gometo.pro in the previous step:
    - for ii = 1:5
    - dset=load(sprintf('2015-05-25_02_%d.txt',ii));
- Copy the active correction fit parameters from cftool to Orchestra
- Record the SoC when rotating Omega; it should be below 1um on all three axes. Use gometo and Matlab or the SoC application for this purpose.

Chi fit parameters
- Run the IDL script chimeto.pro (see 9.3.5) with these parameters:
  IDL> chimeto, SENS=62500, MNUMBER=5, FILENAME='2015-05-25_01' # Note the single quotes!
- Run matlab and generate the coefficients of the fit functions for the Chi rotation:
  - $> cd data
  - $> matlab &
  - Edit the script evaluate.m and adapt the name of the data file and the number of data sets to the configuration applied to chimeto.pro in the previous step:
    - for ii = 1:5
    - dset=load(sprintf('2015-05-25_01_%d.txt',ii));
  - Run the script evaluate.m; this loads the data from the data file into Matlab and displays the deviation graphically in two variants:
    >> evaluate
  - Run the module cftool; this is a curve fitting tool provided by MathWorks; it uses the data set loaded by evaluate.m
    >> cftool
  - Configure cftool:
    - X data: chi
    - Y data: X for a fit on the beamline X-axis; use Y data: Y for a fit on the Y-axis or Y data: Z for a fit on the Z-axis respectively (do one after the other)
    - Use Polynomial for the Chi corrections
    - Select 1 to 5 in Number of Terms:, depending on what achieves the best result without generating useless calculations inside Orchestra: when the RMSE value in the result section of the cftool window doesn't change much when a higher number of terms is selected, then the optimum of calculation efficiency is reached
- Copy the active correction fit parameters from cftool to Orchestra
- Go back to the step Run the IDL script gometo.pro and repeat the procedure until here with minor differences:
  - Run the IDL script chimeto.pro (see 9.3.5) with these parameters:
    IDL> chimeto, SENS=62500, MNUMBER=5, FILENAME='2015-05-25_02'
  - If Matlab is still running, execute the command clear before running evaluate
  - Edit the script evaluate.m and adapt the name of the data file and the number of data sets to the configuration applied to gometo.pro in the previous step:
    - for ii = 1:5
    - dset=load(sprintf('2015-05-25_02_%d.txt',ii));
- Copy the active correction fit parameters from cftool to Orchestra
- Record the SoC when rotating Chi; it should be below 10um on all three axes. Use chimeto and Matlab or the SoC application for this purpose.

# Sub-Procedures
## Mechanical
### Place the Capacitive Distance Sensors
When placing the distance sensors in front of the Prigo, it is irrelevant if the stand is placed upstream or downstream of the goniometer in respect to the beam direction. It is basically also irrelevant in what orientation the aluminum cube that holds the sensors is mounted. What matters is:
1. The assignment of the sensors to the coordinate system in use.
2. If high or low sensitivity is used for the measurements.

1. The coordinate system in use is the right-handed coordinate system of Reference Frame 1, also known as Prigo coordinate system. It was defined by Wayne Glettig and can be described as follows:
- Assuming Omega = 0deg
- The orthogonal axis to the base plate of the Prigo is the Z-axis. It points from Prigo's base to its tip.
- When looking in Z-axis direction, the X-axis points horizontally to the left, the Y-axis points up

Assuming this coordinate system, assign the sensors this way:
- Sensor 1 is the X-axis sensor left of the tip
- Sensor 2 is the Y-axis sensor below the tip
- Sensor 3 is the Z-axis sensor in front of the tip (opposite the Prigo)

If the orientation of the sensors is different (for example the X-axis sensor is right of the tip), then the measurement and calibration applications have to be adapted for sign changes.
- IDL scripts:
  Edit programs (like phimeto.pro) directly and change the signs where necessary
- Java applications:
  Edit the configuration files (like socConfig.xml) and change the sign of the sensitivity value of the affected channel

Align the distance sensors to the Omega rotation axis:
- Mount a steel cylinder with a sphere
- Set sensitivity on all channels to LOW (62500nm/V)
- Attach a scope to all three channels or use the Striptool to display the signal levels
- Rotate the Prigo around the Omega axis; using the adjustment knobs on the linear tables move each sensor to the position where its signal amplitude is symmetrical around 0 when rotating Omega 360deg

2. High or low sensitivity can be selected on the CPL290 units. Depending on this selection, the distance sensors have to be adjusted in the Aluminum block:
- When using low sensitivity, the air gap between the steel sphere and the sensor is 0.65mm. This means that the sensor stands up 13.0mm from the Aluminum block.
- When using high sensitivity, the air gap between the steel sphere and the sensor is 0.05mm. This means that the sensor stands up 13.6mm from the Aluminum block.

Notes:
- The allowed tolerance is in both cases ± 0.1mm.
- The steel sphere has a diameter of 12.7mm.
- See the file PRIGO_Drawings_v.v.pdf for details.

## On the MCS Hand Controller
### Calibrate the linear stage
Navigate to the Calibration Channel menu and select the channel:
MAIN MENU -> SENSOR OPTIONS --> CALIBRATE CHANNEL --> Channel xx

### Search the reference mark on the linear stage
Navigate to the Find Reference Channel menu and select the channel:
MAIN MENU -> SENSOR OPTIONS --> FIND REFERENCE --> Channel xx

### Set the offset of the linear stage to zero
Navigate to the Zero Channel menu and select the channel:
MAIN MENU -> SENSOR OPTIONS --> ZERO CHANNEL --> Channel xx

## On a Console PC
### Check the Prigo applications and EPICS
After starting the prigoTCP service (which starts prigoCONTROL) the applications sometimes don't work properly. The reasons are unknown. However, it can be checked if everything is okay.
Using the medm PRIGO Control Panel, move a tiny bit (say 1mm or 1deg) along SHX, SHY, SHZ, CHI, PHI, OX, OY and OZ and check:
- If the related read back value changes
- If all Smaract stages move (it has been observed that in rare cases Orchestra stopped sending motion commands to one of the stages, reason unknown)

### Disengage the Omega axis
Open the Aerotech rotation stage's expert panel (in the MX lab this is the Aerotech Soloist medm panel with the file name X_PX_soloist_expert.adl) and click on the Disable button

### Engage the Omega axis
Open the Aerotech rotation stage's expert panel (in the MX lab this is the Aerotech Soloist medm panel with the file name X_PX_soloist_expert.adl) and click on the Enable button

### Home the Omega axis
- Open the Aerotech rotation stage's Berger-Lahr panel (in the MX lab this is the Soloist & Berger Lahr medm panel with the file name X_PX_soloist_user.adl) and click on the Home button
- If this doesn't work, check the following items:
  - Berger-Lahr motor driver is switched on
  - The offset in the EPICS motor record is 0 (difference between dial and user positions)

### Run IDL script chimeto.pro
Change the command line parameters passed to chimeto as necessary. See the file header for details.
- $> cd path/to/prigo/git/directory
- $> cd IDL/chimeto
- $> StartIDL
- IDL> chimeto, SENS=62500, MNUMBER=5, FILENAME='2015-06-03_01' # Note the single quotes!

### Run IDL script gometo.pro
Change the command line parameters passed to gometo as necessary. See the file header for details.
- $> cd path/to/prigo/git/directory
- $> cd IDL/gometo
- $> StartIDL
- IDL> gometo, SENS=62500, EFAC=500, MNUMBER=5, VELOCITY=2, FILENAME='2015-06-03_01' # Note the single quotes!

### Run IDL script phimeto.pro
Change the command line parameters passed to phimeto as necessary. See the file header for details.
- $> cd path/to/prigo/git/directory
- $> cd IDL/phimeto
- $> StartIDL
- IDL> phimeto, SENS=62500, EFAC=1000, DATACOLLECTION=0

## On the Prigo PC
### Change and store Prigo IK parameters on the fly
- Login to the prigo3pc (or prigo3pc-backup) as user prigo, open a console window and start OrchestraHMI if not already running: $> OrchestraHMI &
- Connect to the Orchestra instance
- Open the parameter window
- In the navigation bar to the left, select the prigoIK item; this opens a list of changeable parameters
- Select the parameter to change, enter the new value and click on the green tick mark OR set it by draging the slider; when ready click on the Send data to core icon (the two blue arrows pointing downwards), this will make the new value active in the control software
- If the new values should be stored in the configuration file, click on the Save data icon (the diskette icon). This stores the values in this file: /home/prigo/prigoCONTROL/prigoIK/prigoIK.xmp
  Notes:
  - You may want to store the original file under a different name before saving the new values
  - This file will be overwritten when a make install is executed; follow the next step for storing the file in such a way that it will be restored when prigoCONTROL is installed from scratch
- If the new values should survive a make install command from the build environment, store the prigoIK.xmp file in this location:
  - For the production Prigo:
    /home/prigo/prigo/C++/prigo3pc/src/prigoCONTROL/prigoIK/prigoIK.xmp
  - For the backup Prigo:
    /home/prigo/prigo/C++/prigo3pc-backup/src/prigoCONTROL/prigoIK/prigoIK.xmp
Don't forget to check them in to the git repository!
### Set the Prigo IK parameters to their default values
There are various ways how this can be done. A simple way is to copy a prigoIK.xmp file with default values to the relevant location and restart Orchestra. This is described here.
Login to the prigo3pc (or prigo3pc-backup) as user prigo, open a console window and execute these commands:
- $> cd ~/prigoCONTROL/prigoIK
- $> cp ~/prigo/Orchestra/prigo3pc/prigoCONTROL/prigoIK/prigoIK.xmpDflt .
  OR
  $> cp ~/prigo/Orchestra/prigo3pc-backup/prigoCONTROL/prigoIK/prigoIK.xmpDflt .
- $> mv prigoIK.xmp prigoIK.xmpOrig
- $> mv prigoIK.xmpDflt prigoIK.xmp
- $> sudo service prigoTCP restart

Alternatively, the parameters could be set to their default values by using the OrchestraHMI application. This would be online and wouldn't need a restart or the prigoCONTROL and prigoTCP applications.
### Set the Prigo active correction fit parameters to 0
Login to the prigo3pc or prigo3pc-backup PC as user prigo, open a console window and start OrchestraHMI if not already running: $> OrchestraHMI &
- Connect to the Orchestra instance
- Open the parameter window
- In the navigation bar to the left, select the Calibration item; this opens a list of changeable parameters
- Select one parameter after the other, starting from chiXp11 down to omeZc24 and set each one to 0.
- When all are 0, click on the Send data to core icon

### Switch Active Correction on or off
The Active Correction module is (unfortunately) called Calibration in the Orchestra software.
Login to the prigo3pc or prigo3pc-backup PC as user prigo, open a console window and start OrchestraHMI if not already running: $> OrchestraHMI &
- Connect to the Orchestra instance
- Open the parameter window
- In the navigation bar to the left, select the Calibration item; this opens a list of changeable parameters
- Select the parameter P_ENABLED
- If this parameter is set to 1, active correction (the Calibration module) is active. If it is given any other value, active correction is switched off.
- Click on the Send data to core icon

### Start the prigoTCP service
- $> ssh prigo@<Prigo PC>
  - <Prigo PC> is either prigo3pc (for X06DA ES1) or prigo3pc-backup (for X06DA lab)
- $> sudo service prigoTCP start

### Stop the prigoTCP service
- $> ssh prigo@<Prigo PC>
  - <Prigo PC> is either prigo3pc (for X06DA ES1) or prigo3pc-backup (for X06DA lab)
- $> sudo service prigoTCP stop

### Copy the active correction fit parameters from cftool to Orchestra
#### Via OrchestraHMI
Note: Exposing the calibration parameters via the parameter mechanism of Orchestra didn't work. It made the prigoCONTROL application unstable during start-up. Reason unknown.
However, this sub-chapter hasn't been deleted in order to show how the parameters could be changed directly when Orchestra would work as expected.

Omega Correction
If it is the first iteration of the correction, the first digit in the parameter names is always a '1'. If it is however the second iteration, the first digit in the parameter names is always a '2'.
For more information about what parameter from Matlab refers to what parameter in the Prigo see the source file prigo/Orchestra/prigoCONTROL/Calibration/Calibration.cpp.
Example for the first iteration:
- For the X-axis correction:
  - a1 is the parameter omeXa11
  - b1 is the parameter omeXb11
  - c1 is the parameter omeXc11
  - a2 is the parameter omeXa12
  - b2 is the parameter omeXb12
  - c2 is the parameter omeXc12
  - And so on
- For the Y-axis correction:
  - a1 is the parameter omeYa11
  - b1 is the parameter omeYb11
  - c1 is the parameter omeYc11
  - a2 is the parameter omeYa12
  - b2 is the parameter omeYb12
  - c2 is the parameter omeYc12
  - And so on
- For the Z-axis correction:
  - a1 is the parameter omeZa11
  - b1 is the parameter omeZb11
  - c1 is the parameter omeZc11
  - a2 is the parameter omeZa12
  - b2 is the parameter omeZb12
  - c2 is the parameter omeZc12
  - And so on

Chi Correction
If it is the first iteration of the correction, the first digit in the parameter names is always a '1'. If it is however the second iteration, the first digit in the parameter names is always a '2'.
For more information about what parameter from Matlab refers to what parameter in the Prigo see the source file prigo/Orchestra/prigoCONTROL/Calibration/Calibration.cpp.
Example for the first iteration:
- For the X-axis correction with a 3rd order polynom:
  - p1 is the parameter chiXp13 (coefficient for x^3)
  - p2 is the parameter chiXp14 (coefficient for x^2)
  - p3 is the parameter chiXp15 (coefficient for x^1)
  - p4 is unused
- For the Y-axis correction:
  - p1 is the parameter chiYp13 (coefficient for x^3)
  - p2 is the parameter chiYp14 (coefficient for x^2)
  - p3 is the parameter chiYp15 (coefficient for x^1)
  - p4 is unused
- For the Z-axis correction:
  - p1 is the parameter chiZp13 (coefficient for x^3)
  - p2 is the parameter chiZp14 (coefficient for x^2)
  - p3 is the parameter chiZp15 (coefficient for x^1)
  - p4 is unused

#### Via Calibration.cpp and a rebuild
- Disconnect OrchestraHMI from the running instance
- Stop the prigoTCP service
- Copy the parameters from the cftool window to Calibration.cpp
- Save Calibration.cpp
- Rebuild the project and copy the executable module to the installation directory:
  - $> cd path/to/prigo/git/directory
  - $> cd Orchestra/prigo3pc-backup
  - $> make build
  - $> cp prigoCONTROL/Calibration/Calibration.dlc ~/prigoCONTROL/Calibration/
- Start the prigoTCP service
- Using the medm PRIGO Control Panel:
  - Connect to the IOC
  - Initialize the Prigo's axes (move the distance sensors away beforehand, and back into position afterwards; the Z-axis is well suited for this purpose)
  - Move SHX, SHY, SHZ, OX, OY, OZ to their set positions. Notes:
    - After the reboot, the actual position is a default position, which is different from the position after the execution of phimeto.pro and gometo.pro.
    - To move to the set position, click on the value in a blue field (don't change it!) and press <Enter>.
  - Check the Prigo applications and EPICS
- Optional: connect OrchestraHMI from the running instance; check if active correction is on (see 9.4.4)

# Glossary
IK	Inverse Kinematics
MCS	Modular Control System (term used by SmarAct)

## Explanation of Terms
### Aerotech base stage
An Aerotech base stage usually consists of a linear X/Y/Z base stage and a vertically mounted air-bearing rotary stage. The Prigo is mounted on the front of the rotary stage.
In a lab, a base stage may only consist of a vertically mounted rotary stage.

# External References
- Backup Prigo in the new MX Wiki:
  https://intranet.psi.ch/SLS_MX/Goniometer06D
- Sphere of Confusion application in the new MX Wiki:
  https://intranet.psi.ch/SLS_MX/SocApplication
- Production Prigo in the old MX Wiki:
  https://intranet.psi.ch/MXGroup/PRIGO
- Git repository of Prigo code (Documents, IDL, Java, Matlab, Orchestra, Smaract):
  https://github.psi.ch/projects/MXEC/repos/prigo/browse
- See Git repository, Smaract folder